## Sheet1
Example calculation of Cr concentration in Cr-augite from raw peak and background intensities
See the EPMA K-ratio and wt% calculation pdf for more info
The peak intensity is measured by driving the wavelength spectromer to the Cr K-alpha peak using the LiF diffracting crystal
The background intensities traditionally were determined by driving to two offset positions to measure a high and low offset background, then interpolate to obtain the background at the peak position
The k-ratio is then calculated from k = (P-B)smp / (P-B)std using the measured and stored values for the peak and background intensities on the sample and standard
The intensities used are in cps/nA to correct for time and probe current differences between sample and standard measurements (they are also corrected for drift in X-ray intensities and probe current)
So 10000 counts per 100 seconds, measured at 25nA is 10000/10/25 = 100 counts per second = 4 cps/nA
This example shows the peak - background intensity values in cps/nA
k-raw is the relative intensity on the sample compared to the working standard (typically not a pure element); k-raw = 1 means the same intensity is measured on the sample as on the standard
We use working standards that are high in concentration but similar enough to the samples for valid calculation of the matrix correction. This yields a measurement with high counts per unit wt% of the element in the standard.
Counting statistics considerations show that high count rates on the standard propagate to increased confidence of the measurement on samples
Here the primary standard for Cr is Cr2O3 glass, with emitted intensity for Cr K-alpha of 424.21 cps/nA
The k-raw value indicates that 0.77% of the intensity is emitted from the Cr-augite pyroxene compared to the Cr2O3 standard
The ZAF matrix correction factors for the Cr-augite are calculated using all elements in the augite, and is calculated iteratively as progressive estimates of the concentration have been corrected.
For Cr in the Cr-augite, the ZAF factors are: Z = 1.1746, A = 1.0144, and F = 0.9894, multiplicative combination ZAF = 1.1789
The ZAF factors and standard correction factor can be calculated using the CalcZAF program; this is done auto-magically for you using the Probe for EPMA software when you click on the Analyze button
|  | Zedcor | Abscor | Flucor | ZAFCOR
| Cr in Cr-augite | 1.1746 | 1.0144 | 0.9894 | 1.1789
Cr correction factor for Cr2O3 = 0.6395; this is the emitted intenstiy for Cr K-alpha from Cr2O3 compared to pure Cr metal, and is calculated by running the ZAF correction in reverse
The calculation is C = k * ZAF * CF, or just C = k * ZAF if the standard is a pure element
If C is known, then the k-ratio and ZAF factors are calculated from k = C / ZAF and ZAF = C / k
C is entered as weight fraction for these calculations and converted to element wt% by multiplying by 100
Note that we measure the elements in a material and calculate associated oxygen by stoichiometry
Here the C of Cr in Cr-augite is k-raw * ZAF * CF * 100 to yield element wt% Cr
The accepted value for Cr in the augite is 0.582 wt %, so the result is excellent with observed / accepted = 1.0016
We usually round element concentration to 2 decimal places but this depends on the counting statistics
For quality control in our microprobe runs, we analyze Kakanui hornblende as a true secondary standard and use it to identify systemic, spectrometer, and element-specific problems
|  | P-B sample | P-B standard | Relative intensity
|  | smp | std | kraw | ZAF | Std cf | C | C wt% | Accepted | obs/acc
| Cr in Cr-augite | 3.28 | 424.21 | 0.007732 | 1.1789 | 0.6395 | 0.0058292 | 0.58292 | 0.582 | 1.0016